## Slide 1: What are we measuring with M/EEG?
✉ g.o’neill@ucl.ac.uk @ofluff SPM M/EEG Course
George O’Neill WCHN, Queen Square Institute of Neurology University College London

## Slide 2: Overview
The biophysical origin of the M/EEG signal
M/EEG instrumentation
Source sensitivity
An brief introduction to source reconstruction: forward modelling

## Slide 3: Part I
Biophysical origin of the M/EEG signals
M/EEG instrumentation.
Source sensitivity.
A brief introduction to source reconstruction: forward modelling

## Slide 4: EEG and MEG
Electroencephalography (EEG) and magnetoencephalography (MEG) are non-invasive means that measure the same underlying neural currents.
EEG: Measures differences in electric potentials on the scalp
MEG: Measures changes in magnetic flux density outside of the head

## Slide 5: Origin of the MEEG signal
Cortex
CSF
Skull
Scalp
White matter
Spocter et al, (2017)

## Slide 6: Origin of the MEEG signal
Spocter et al, (2017)
Pyramidal cells
Found in layers II/III and V
Organised so primary axis is perpendicular to the cortical surface
Open field layout makes this detectable with M/EEG
Stellate cells
This example from layer IV
Closed fields cancel out, not seen by M/EEG

## Slide 7: Post Synaptic Potentials
Synaptic input leads to ionic currents across the postsynaptic membrane
Excitatory Post-Synaptic Potential (EPSP): influx of positive Na+ ions at apical dendrites causes depolarization of the postsynaptic cell
Extracellular volume currents complete the loop of ionic flow so that there is no build-up of charge
MEG is more sensitive to intracellular currents, EEG to extracellular

## Slide 8: Post Synaptic Potentials
Synaptic input leads to ionic currents across the postsynaptic membrane
Inhibitory Post-Synaptic Potential (IPSP): influx of negative Cl- ions causes hyperpolarization of the postsynaptic cell

## Slide 9: What about action potentials?
Action potentials (APs) are unlikely to contribute to the M/EEG signal.
Action Currents
Postsynaptic Currents
Synapse
B
Jvol
APs produce electric quadropoles, which their intensity decline steeply with distance (1/r3). Postsynaptic currents are dipolar, which drop off as 1/r2.
APs have a short duration. Would need to be highly synchronised to be measurable.
The bi-phasic nature of the depolarising and repolarising currents might result in mean field cancellation.

## Slide 10: How many neurons?
A single neuron is not detectable.
Neuronal models of detailed morphology were simulated, excited by virtually injecting current.
Equivalent current dipole (ECD) moment was estimate by summing across dipoles in an area.
50000 cells sufficient to generate a dipole source of 10 nAm
Murakami et al (2006)

## Slide 11: Primary and volume currents
MEG Coils
EEG Electrodes
Magnetic field
Volume currents
MEG measures changes in the magnetic field generated by an electric current (Sarvas 1987, Hämäläinen et al. 1993)
These magnetic field are mainly induced by primary currents based on excitatory activity (Okada et al. 1997).
Volume currents yield potential differences on the scalp that can be measured by EEG.
Careful!
Volume current fields are heavily distorted by the insulating properties of the skull. This can limit the spatial resolution of EEG.

## Slide 12: Part II
M/EEG instrumentation.
Source sensitivity.
A brief introduction to source reconstruction: forward modelling
The biophysical origin of the M/EEG signal

## Slide 13: EEG Signal Detection: Electrodes
EEG records potential differences at the scalp using a set of active electrodes and a reference (ground) electrode.
The reference electrode is important as it eliminates noise from the amplifier circuit
The representation of the EEG channels is referred to as a montage
  Unipolar/Referential  potential difference between an electrode and a designated reference
  Bipolar  represents the difference between a pair of adjacent electrodes (e.g. ECG, EOG, EMG)
Potential differences are amplified, filtered and digitized.
McFarland et al., (1997)

## Slide 14: EEG Signal Detection: Electrodes
EEG records potential differences at the scalp using a set of active electrodes and a reference (ground) electrode.
The reference electrode is important as it eliminates noise from the amplifier circuit
The representation of the EEG channels is referred to as a montage
  Unipolar/Referential  potential difference between an electrode and a designated reference
  Bipolar  represents the difference between a pair of adjacent electrodes (e.g. ECG, EOG, EMG)
Potential differences are amplified, filtered and digitized.
McFarland et al., (1997)

## Slide 15: MEG Signal Detection: SQUIDs
Superconducting QUantum Interference Device
An ultra sensitive detector of magnetic flux (Φ). Consisting of two superconductors with “insulating” layers, forming two Josephson Junctions in parallel.
The measured voltage across the junctions is dependent on magnetic flux.
Measures changes on order of femtoTeslas (10-15 T).
Φ0 = 2.06783383 × 10-15 Wb
(The magnetic flux quantum)
A 1cm2 junction has a dynamic range of 10-11 T

## Slide 16: MEG Signal Detection: SQUIDs
Superconductors only become superconductors below in a certain temperature. The SQUIDs in a MEG are kept at 4 K (-269.15 °C) using a bath of liquid helium.
This makes SQUID-MEG quite limited in its applications.
Participant has to keep head still relative to sensors!
Bulky instrumentation
Liquid helium is expensive.

## Slide 17: MEG Signal Detection: OPMs
A new generation of sensors which are just a sensitive, but can operate safely on the scalp. The most commonplace so far are Optically Pumped Magnetometers.
Theoretical improvements in SNR over SQUID MEG. EEG form factor a bonus.
Boto et al. (2018)
QuSpin Inc.
Warning No turn-key OP-MEG solution exits yet. Proceed with caution!
WCHN 2021

## Slide 18: MEG Signal Challenges
The MEG signal is tiny!
We need approaches to shield our sensors from this noise to be able to measure anything useful.
Vrba (1999)

## Slide 19: Magnetically shielded rooms
MEG systems are currently housed within Magnetically Shielded Rooms (MSRs), which give passive shielding against noise from the environment. The best rooms can get the residual background fields to the order of a few nanoTeslas.
Concentric shells of mu metal, copper and aluminium bend external fields around the MSR
Shielded interference

## Slide 20: Magnetically shielded rooms
Shielded interference
Some external interference is time-evolving. The MEG at UCL is ~50 m from passing trains, which the MSR cannot eliminate in its entirety.
Holmes et al. (2019)
Trains departing nearby

## Slide 21: From magnetometers to gradiometers
Magnetometer
Axial Gradiometer
Gradiometers used a second counter-wound coil a given distance aware from the primary coil, which allows it to sense the gradient of a field between the two coils.
Planar Gradiometer

## Slide 22: From magnetometers to gradiometers
Axial Gradiometer
Measured field
Signal
Noise
Magnetometer
Gradiometer
Sources with which generate large magnetic fields but are far away produce small field gradients.
Nearby sources generate much larger gradients.
This allows us to push up the SNR of our data.

## Slide 23: From magnetometers to gradiometers
The field patterns from axial gradiometers and planar gradiometers look quite different.
fieldtriptoolbox.org
Axial
Planar

## Slide 24: Part III
Source sensitivity.
A brief introduction to source reconstruction: forward modelling
The biophysical origin of the M/EEG signal
M/EEG instrumentation

## Slide 25: Source orientation
Scalp
Skull
Cortex
Bulk current flow from EPSPs is oriented perpendicular to the cortical ribbon.
Sulcus
Gyrus
Orientation of source is important….
EEG can detect both tangential and radial oriented dipoles
MEG can see tangential, but struggles to sense radial dipoles
  If you assume a symmetric sphere for the head, radial dipoles are impossible to detect!

## Slide 26: Source orientation
So what’s the point if MEG cannot see sources in the gyri/sulci?
Hillebrand and Barnes (2002)
The head isn’t a perfect sphere.
Even if you assume it is very few sources are truly radial
Most sources are not point sources, so will likely contain a tangential component.
tl;dr - it’s not a massive concern.

## Slide 27: Source extent
Ahlfors et al. (2010)
Extended and neighbouring sources can lead to source cancellation at the sensor level.

## Slide 28: Source depth
Depth is a limiting factor in MEG measurements.
Hillebrand and Barnes (2002)
Sensor-level amplitude decreases with distance from source (1/r2)
Deeper sources appear more radial, which MEG is less sensitive to

## Slide 29: Source depth
Its hard, but not impossible given sufficient SNR
Pizzo et al. (2019)
Meyer et al. (2017)

## Slide 30: Source depth
If SNR is good enough, you should be able to discriminate between layers within the cortex!
Bonaiuto, Little, et al. (2021)
Beta bursts are hypothesised to employ both layer II/III and V pyramidal cells
MEG can pick apart which part of the event occurs for which layer.

## Slide 31: Part IV
A brief introduction to source reconstruction: forward modelling
The biophysical origin of the M/EEG signal
M/EEG instrumentation
Source sensitivity.

## Slide 32: Source reconstruction
From sensors…
…to sources!
Forward problem
Inverse problem

## Slide 33: A generative model of sources
Starting from our sensor-level M/EEG data, y, for a given time, t, the we can assume is generated by sources in the brain:
Lead fields:
How a source is represented at sensors
Source strength at time t
Everything unexplained (noise)
Encapsulate all time and we can turn this into a set of matrices.
All data across all time
All lead fields
All source activity across all time
‘noise’ across all time

## Slide 34: Lead fields: the forward problem
Lead fields
The forward problem
If we know the precise position/orientation/amplitude of a dipole in the brain, can we estimate what the associated senor-level pattern should be?
YES!
For a given dipole, there exists one unique solution. This makes it straightforward to solve!
For MEG data, the solutions are typically a lot simpler than for MEG data.

## Slide 35: Lead fields: the mathematical formulation
Sorry. But this is for your own good.
B
JP
JV
The measured current J, which is in a volume G, with conductivity σ, at a point in space r’, is the sum of the (intracellular) primary currents and (extracellular) volume currents.
The magnetic field B produced from J as seen from a location r, follows the Biot-Savart law:
Electric Potential

## Slide 36: Fields in an infinite medium
In an infinite medium with constant conductivity :
The volume currents do not contribute.
We can simplify further by assume
  is the direction and strength of a current dipole
  at location , 0 everywhere else
Q
r-rQ
Sensor
Primary Magnetic Field
Primary Electric Potential

## Slide 37: Fields in an inhomogeneous medium
The infinite model, whilst a nice toy is too simplistic… Assuming the head is set of domains with constant conductivity within each encapsulated by a surface S…
σ1
σ2
S1
S2
n1
n2

## Slide 38: What?
I’ve probably lost you, so lets break this down what this is telling us.
The observable magnetic fields are partly dependent on the volume currents, so need to consider potentials.
To calculate the potentials outside of the head, you need to first numerically solve potentials on the boundaries between each regions.
Anatomical properties needs to be considered!
This could be incredibly tedious…

## Slide 39: Forward models
Various approximations to the problem are available.
Single Sphere
Local Spheres
Corrected Sphere (a.k.a. single shell)
Boundary Element Model (BEM)
Finite Element Model (FEM)
Simple
Complex
EEG requires more complex models which are able to predict electric potential differences better than the simple spherical models (typically a 3-shell BEM or better).
This does require more knowledge about the anatomy.

## Slide 40: Forward models
MEG may benefit from more complex models and accurate segmentations.
Stenroos et al. NIMG (2014)
Increased segmentation accuracy
Simpler models
Reference model

## Slide 41: Forward models
Which model do I recommend?
Note: this is my personal opinion.
For most users…
EEG: 3-Shell Boundary Element Model (BEM) (or better)
MEG: Nolte’s Correct Sphere (a.k.a. Single Shell) (or better)

## Slide 42: The Inverse problem
The inverse problem
Given a set of observations at the sensor level, can we estimate where in the brain they originate from?
Can we solve for J by doing….
No!
This problem is ill-posed for a number of reasons.
There are many more sources than sensors
Opposing neighbouring dipoles can cancel out, leading to an infinite number of possible solutions…
How we solve for J in these limitations will be discussed later in the course!
Source Activity

## Slide 43: Summary
MEG and EEG are direct measures of neural activity, with high temporal resolution.
This signal is predominantly based on the ensemble postsynaptic currents of 10s of thousands of pyramidal cells.
MEG is biggest challenge is ensuring that the SNR is high enough given interference several orders of magnitude larger than the signal.
EEG has higher sensitivity to both deep and radial sources, but the signal is strongly affected by conductivity changes between brain, scull and scalp. Which makes spatial precision a challenge.
Forward models describe how the currents in the brain give rise to electric potentials or magnetic fields measured outside the head

## Slide 44: Thanks for listening!
Any questions?

## Slide 45: Lead fields: the forward solution
The forward solutions describes the sensor activity you’d expect to see from a unit strength source given a location and orientation.
The leadfield matrix L depends on:
sensors type, location and orientation
source space, (is it a cortical surface or volumetric grid?)
head geometry
conductivity of the head tissues